SMP15A3 PRESSURE SENSOR	REV 1.00

# Characteristics	8/06/02

All parameters measured at 5 volts excitation at ambient temperature 25℃, unless otherwise specified.

| Parameter | Min | Typ | Max | Units |
| --- | --- | --- | --- | --- |
| Excitation Voltage | 0 | 5.0 | － | Volts |
| Excitation Current | 0 | 1.0 | － | mA |
| Span | 70 | 100 | 130 | mV |
| Zero Offset | -10 |  | +10 | mV/V |
| TC Span | － | -0.22 | － | %FS/℃ |
| TC Offset | -0.08 |  | +0.08 | %FS/℃ |
| Thermal Hysteresis | -0.1 |  | +0.1 | %FS |
| Linearity | -0.3 |  | +0.3 | %FS |
| Bridge Impedance | 4.0 | － | 6.0 | K |
| Proof Pressure | 2X |  |  | Rated FS Pressure |
| Operating Temp | -20 |  | 85 | C |
| Storage Temp | -40 |  | 125 | C |

NOTES

- TCO and TCS are guaranteed by design. Therefore these parameters are not 100 % tested.

- Linearity is the maximum deviation of measured output from “best straight line”.

- The wetted material is nickel and silicone gel, and is intended to use with non-corrosive gaseous media.